## Read Me
Biocompatibles Toric Lens Calculator
This spreadsheet was developed by Geoff Cooke, Tim Green and Graeme Young and was originally described in an article in the Optician (Cooke G, Green T, Young G. Spreadsheet calculation of oblique cylinder powers. Optician 1996; 211(5540): 28-30). The original formulae quoted in that article contained some errors and these have been corrected in this version.
Toric lens rotation | When toric soft lenses fail to give satisfactory visual acuity (VA) due to misorientation, several signs help to confirm whether the reduced VA is due to simple misorientation rather than inaccurate refraction, inaccurate lens manufacture or a combination of factors: (1) the over-refraction will yield a resultant cylinder power which is twice the resultant sphere power; (2) when the lens is digitally rotated to the intended orientation, the VA will improve to the expected level. In these cases, the remedy is to reorder the lens taking into account the axis mislocation.
Other causes of reduced VA | When poor vision cannot be explained by rotation, it is helpful to verify the lens specification and possible re-check the patient’s refraction and any back vertex distance calculation. Usually, these investigations will yield the source of the error and the practitioner can reorder a new lens based on his or her findings. Occasionally, however, an alternative and possibly easier option is to reorder the new lens based on a calculation of the resultant of the lens power (measured rather than nominal) and an accurate over-refraction. This is sometimes appropriate with high prescriptions or in cases where the findings are inconsistent and a solution is not obvious.
Lens Accuracy | Remember that the tolerances for toric soft lenses are wide and that it is therefore unwise to re-order lenses based on the over-refraction and the labelled power.
It is helpful to verify the lens specification by blotting with a medical wipe and measuring the lens on a focimeter. The cylinder axis can be read more accurately by using a Tori-check lens holder. In some instances, it is also worth re-checking the patient’s refraction and any back vertex distance calculation.
When to use the calculator | Reordering a new lens based on a calculation of the resultant of the lens power (measured rather than nominal) and an accurate over-refraction is sometimes appropriate, e.g. with high prescriptions, in cases where the findings are inconsistent, or - as a last resort.
Using the calculator | Insert the lens power (rows 3-5) and over-refraction details (rows 6-8) in the appropriate cells. Also note the amount of lens rotation (in degrees) in the cell on row 9. To indicate the direction of lens rotation, clockwise rotation has a plus sign and anti-clockwise rotation a minus sign.
Rounding | The outputs can be read either to two decimal places (columns B&C) or rounded to the nearest 0.25D (columns E&F).
Expanding the calculator | The number of columns can be increased by highlighting a column, selecting the square in the top right hand corner and dragging across other columns.
Disclaimer | This spreadsheet is intended as an aid to toric lens fitting and supplements rather than replaces the skill of the practitioner. No liability will be accepted for problems arising from the use of this spreadsheet.
Graeme Young Nov. 2001

## Oblique cyls
Toric Lens Calculator
Inputs - Original Lens | R | L | R | L
Lens sphere power | -3.5 | 0 | -1 | -0.75
Lens cylinder power | 1.25 | -2 | -2 | -1
Lens axis | 180 | 180 | 180 | 170
Sphere power of over-refraction | 0.5 | 0 | 0.5 | -0.75
Cylinder power of over-refraction | -1 | -2 | -1 | -1
Axis of over-refraction | 20 | 180 | 20 | 170
On eye lens rotation (+ =CW) | 5 | -15 | 5 | -10
Calculations
CO total | 0.46497 | -3.7321 | -2.7357 | -1.9397
C45 total | -0.85985 | -1 | -0.29549 | 0.34202
C1 | -0.97751 | -3.8637 | -2.7516 | -1.9696
A1 | 64.2012 | -7.5 | 8.0824 | -15
Outputs - New Prescription
Sphere power | -2.3862 | -0.068148 | -0.5 | -1.5
Cylinder power | -0.97751 | -3.8637 | -2.75 | -2
Axis | 64 | 172 | 8 | 165